|  | Nordisk Sikkerhet AS |
| --- | --- |
|  | Tender title: Supply of vehicle for radiation monotoring and post-detection activities |
|  | Project title: “Post-detection equipment and training for Lisichansk unit of Luhansk border detachment, SBGS, Ukraine” |

Specifications

Supply of vehicle for radiation monitoring                              and post-detection activities

(mobile radiological complex)

Contracting Authority: Nordisk Sikkerhet AS

Recipient: Lisichansk unit of Luhansk border detachment of SBGS

Tenderer’s name: __________________________________

Contents

1.	VEHICLE FOR RADIATION MONITORING AND POST-DETECTION ACTIVITIES (MOBILE RADIOLOGICAL COMPLEX)	3

2.	LABORATORY COMPARTMENT	7

3.	DECONTAMINATION EQUIPMENT AND DETERGENTS	9

4.	PERSONAL PROTECTIVE EQUIPMENT	9

5.	SYSTEM FOR AUTOMATED RADATION MONITORING OF THE ENVIRONMENT (INCL. GROUND GAMMA SERVEY)	12

6.	WEATHER STATION	9

7.	ISOTOPE IDENTFIFICATION DEVICE (with gamma and neutron channels)	16

8.	MULTI-PURPOSE SURVEY METER (with alpha, beta, gamma and neutron channels)	18

9.       SET FOR COLLECTING SAMPLES………………………………………………………………………22

10.	BACKPACK FOR RADIONUCLIDE IDENTIFICATION AND SOURCE LOCATION	22

11.	PERSONAL ELECTRONIC DOSIMETERS	24

12.	AUXILARY EQUIPMENT FOR POST-DETECTION ACTIVITIES	24

13.	VISIBILITY OF NORVEGIAN FINANCING	24

14.	DOCUMENTATION	27

15.	DELIVERY TERMS AND CONDITIONS	28

16.	TRAINING COURSE: OPERATION, MAINTENANCE AND REPAIR	28

17.	WARRANTY AND POST-WARRANTY SERVICES	29

18.	TIME SCHEDULE	30

NOTE

The Tenderer shall fill in the Annex “Specifications” in the format given below.  The Tenderer’s proposed supplies should be manufactured and certified in accordance with the technical regulations and standards of Ukraine and/or EU countries. The complete table should be submitted to the Contracting Authority along with the required tender documents. On the front page of the Annex “Specifications”, the Tenderer shall indicate its name. After the completion of this document, it should be signed and dated by the Tenderer-authorized person.

# VEHICLE FOR RADIATION MONITORING AND POST-DETECTION ACTIVITIES (MOBILE RADIOLOGICAL COMPLEX)

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 1 complex |  |
| TECHNICAL SPECIFICATIONS |
| Type | Mobile radiological complex based on a motor vehicle (motor vehicle is included in the scope of this contract) |  |
| Functions | The Mobile radiological complex shall be equipped with measuring instruments and tools needed for: (a) Detection of RN materials, (b) Quick search, characterization and radionuclide analysis of the revealed RN material, (c) Radiological monitoring on broad territories, (d) Individual radiation dose monitoring for expert response team, (e) Post-detection activities such as: (i) Decontamination of equipment/vehicle (small areas), (ii) Decontamination of road surfaces and terrains (small areas), (iii) Personnel decontamination. |  |
| Equipment | The Mobile radiological complex shall be equipped with the following: - Vehicle, - Automated workplaces for two operators (equipped with two tables, two seats, two PCs and multifunction printer), - Radiocommunication equipment (stationary radio station with five walkie-talkie sets with headphones), - Automated radiation monitoring system, portable measuring instruments and set for collecting samples, - Weather station, - Decontamination equipment and detergents for use against radioactive contamination (incl. two small portable decontamination systems for manual use and three handheld pressure sprinklers for detergents), - Personal protective equipment and personal decontamination kits. |  |
| Auxiliary equipment | The Mobile radiological complex shall be equipped with the following: - Set of lights (searchlight and a tower for it, hand lamp, emergency light, head light, etc.), - Warning signs and safety tapes to mark the boundaries (areas) of contamination with toxic and radioactive substances, - Devices for automatic placement of safety tapes and warning signs (one stationary and portable ones), - Photo/video cameras to ensure gathering of evidences and incident investigation. |  |
| Placement of equipment | All the components should have a modular design and be firmly fixed inside the vehicle |  |
| Power supply | From mains of AC current, 220 ±10% V, 50±5 Hz and from generator |  |
| Service life of vehicle, no less than | 10 years |  |
| ENVIRONMENTAL REQUIREMENTS TO OPERATION AND STORAGE |
| Ambient temperatures | From -30 to +45 0C |  |
| Relative humidity | Up to 95% at ambient temperature of 350C and lower, without condensation of moisture |  |
| VEHICLE |
| Type of vehicle | Van |  |
| Class of vehicle | Class B |  |
| Body type | Van |  |
| Load carrying capacity | No less than 1500 kg. Higher load carrying capacity is an advantage |  |
| Colour | Dark green is recommended |  |
| Drive gear | 4-wheel drive |  |
| Engine | Turbocharged direct injection (TDI), Power of 100 kW or greater, Volume of 2000 cm3 or greater. |  |
| Type of gearbox | Mechanical, 6 gears or better |  |
| Fuel type | Diesel |  |
| Number of seats | First-row seats: 3 seats, including driver. Laboratory compartment: 2 seats, seats-transformers (360º turnover) are an advantage |  |
| Sound, vibration and thermal insulation | Yes, both for the laboratory and rear compartments |  |
| System assisting the cold-start of diesel engine | Yes |  |
| VEHICLE CONFIGURATION |
| Parking radar | Yes, parking radar system with rear video camera |  |
| GPS navigator | Yes |  |
| GPS monitoring system | Yes, able to communicate the position data to the head office of Luhansk border detachment |  |
| Video surveillance system | Yes, 360º video surveillance system with video recorder |  |
| Additional set of winter tires and car wheels | Yes |  |

| Climate control system (or air conditioning) | Yes |  |
| --- | --- | --- |
| Additional heater | Yes, feeding from the fuel tank of the vehicle |  |
| Radio station | - Stationary radio station – 1 set; - Portable radio stations (walkie-talkie) – 5 pcs.; - Active headphones with headset for communication inside the expert response group – 5 pcs. |  |
| Hydraulic jack | Yes, of 5 t |  |
| Winch | Yes, two winches: - stationary one, on the front reinforced bumper, - portable one, located in the rear compartment. |  |
| Tow bar | Yes |  |
| Car tool kit | Yes |  |
| Entrenching tools | Yes, 1 set including the big shovel, small shovel, axe, saw, pail |  |
| Tow rope of the appropriate tonnage | Yes |  |
| Starting cables | Yes |  |
| Anti-roll blocks | Yes, 2 sets (2 pcs. in one set). |  |
| Chains on wheels (for movement in deep snow) | Yes |  |
| Car wheel anti-skid pads | Yes, 2 pcs |  |
| First aid kit | Yes |  |
| Reflective vests | Yes, 4 pcs. |  |
| Fire extinguisher | Yes, 3 pcs of 5 kg (one per each compartment) |  |
| Rooftop luggage rack | Yes, removable luggage rack |  |
| Backdoor luggage rack | Yes, removable luggage rack |  |
| Road protection | Yes, armor protection plates for engine and chassis. |  |
| Raised air intake system | Yes, adapted for overcoming water obstacles (i.e.,raised air intake system/snorkel) |  |
| Additional external lights | Yes, i.e. boot-mounted spotlights to illuminate the area behind the vehicle |  |
| Transformer ladder | Yes, having 4 sections |  |
| Invertor for the AC mains | Yes, converter 12V DC to 220V AC |  |
| Portable generator | Yes, portable (small) diesel generator, AC current 220V, power of 2,5 kW or greater, feeding from the fuel tank of the vehicle |  |
| Additional accumulator battery | Yes, of 12 V, 2 Gel batteries x 140 Ah |  |
| AIR CLEANING SYSTEM |
| Air cleaning system | Yes, air filtration system based on 1 (m filters against airborne radioactivity (radioactive aerosols) |  |
| Air pressure | Positive air pressure in the driver’s cabinet and laboratory compartment |  |
| Filters | Yes, 10 filter cartridges for radioactive aerosols from 1 (m |  |
| OTHER REQUIREMENTS |
| Certified for use | Yes, the equipment of the Mobile radiological complex shall be certified for use in Ukraine and/or in EU in accordance with the Ukrainian Legislation in force |  |
| Dual-use items (goods, software or technologies) | Not accepted. The Mobile radiological complex shall not be subject to export control, shall not contain dual-use goods, software or technologies and must not be under export/import restrictions of a similar nature. |  |

# LABORATORY COMPARTMENT

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| LABORATORY COMPARTMENT |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 1. Operator’s desks – 2 pcs, 2. Operator’s seats – 2 pcs, 3. PCs – 2 pcs, 4. Multifunction printer – 1 pc, 5. Cupboard and rack with shelves – 1 set. |  |

| TECHNICAL SPECIFICATIONS |
| --- |
| Wall between the driver’s cabinet and the laboratory compartment | Yes |  |
| Wall between the laboratory and the rear compartment | Yes |  |
| Number of automated operator’s workplaces | 2 |  |
| Operator’s desk | Yes, for 2 operators minimum |  |
| Operator’s seat | Yes, for 2 operators minimum. Seats-transformers (360º turnover) are an advantage. |  |
| Safety belts at operator’s workplace | 2 minimum. All the seats shall be secured with seatbelts, so that it is possible to occupy them by operators while vehicle is in motion. |  |
| Cupboard and rack with shelves | Yes, cupboard and rack with shelves to store and charge equipment and accessories, including fasteners to fix the instruments on the shelves and shock protection of the equipment. Cupboard and rack should contain the identification labels for equipment and accessories. |  |
| Operation from external power supply | AC current, 220 ±10% V, 50±5 Hz |  |
| Multifunction printer | Yes, multifunction device for printing, scanning, copying documents (MFD), having the following or better characteristics: - Laser black-and-white printing, - Paper size of A5 to A4. |  |
| Automated operator workplace | On the base of the PC having the following or better characteristics: - Shockproof system unit (or laptop), - Quadra-core CPU 2.5 GHz Core i5 or greater; - RAM 16 GB; - HDD/SSD 256 GB; - External connections: USB 3.0 or higher ≥ 3; - Monitor of 15.6" or larger; - Wireless keyboards; - Wireless mouse. |  |

| Shockproof | Yes, except for multifunction printer |  |
| --- | --- | --- |
| Waterproof | Yes, IP54 or better, except for multifunction printer |  |
| System and application software for automated operator workplace | Yes, including Win 10 Pro (licensed CD-ROM shall be included in the scope of supply) |  |
| Software licenses | Yes, for 5 years of operation |  |

# DECONTAMINATION EQUIPMENT AND DETERGENTS

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| DECONTAMINATION EQUIPMENT |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 1. Small portable decontamination system for manual use – 2 pcs, 2. Handheld pressure sprinkler for detergents – 3 pcs., 3. Detergents – 90 kg, 4. Personal decontamination kits – 30 pcs. |  |
| Purpose | (i) Manual decontamination of people, (ii) Manual decontamination of equipment/vehicles (small areas), (iii) Manual decontamination of road surfaces and terrains (small areas) against radioactive agents |  |
| Canister capacity, no less than | 10 liters for small portable decontamination system, 1 liters for handheld pressure sprinkler for detergents. |  |
| Accessories | Set of brushes, Gun for detergents, Telescopic pole (of 1 m long or more) |  |

| DETERGENTS |
| --- |
| Applied detergents (to be included in the scope of supply) | Mobile radiological complex shall allow the use of the following solution: 25 % sulphanole, 50% three polyphosphate, 18% sodium sulfate or equivalent decontaminants for equipment and people. |  |
| Quantity of detergents (to be included in the scope of supply) | 90 kg of detergent |  |
| Resistance to decontaminating agents | Surfaces which may come into contact with radioactive materials shall be manufactured from corrosion-resistant materials or have a surface coating (paint) which is resistant to decontaminating products |  |

# PERSONAL PROTECTIVE EQUIPMENT

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 1. Protective suits - 10 pcs, 2. Protective gasmasks - 10 pcs, 3. Filtering and absorbing elements – 20 pcs. |  |
| PROTECTIVE SUITS |
| Radiation protection | Protection against surface radiation contamination (radioactive particles) |  |
| Category | Category III as per EU standard. Chemical protection clothing |  |
| Components | Overall (10 pcs), Protective gloves (10 pairs), Protective boots (10 pairs), Protective overboots (100 pairs), Carrying bag (10 pcs). |  |

| Material | TYCHEM 6000F or equivalent |  |
| --- | --- | --- |
| Material of boots/high boots | Butyl rubber |  |
| Material of gloves | Butyl rubber |  |
| Protection against radioactive particles | EN 1073-2:2002(1) |  |
| Protection against liquids under pressure of 3 Bar | Type 3B - EN 14605:2005 (DSTU EN 14605:2007) |  |
| Protection against liquid aerosols | Type 4 - EN 14605:2005 (DSTU EN 14605:2007) |  |
| Protection against airborne solid particles | Type 5 - EN ISO 13982-1:2004 (DSTU EN ISO 13982-1:2009) |  |
| Protection against liquid chemicals | Type 6 - EN 13034:2005 (DSTU EN 13034:2007) |  |
| Biological barrier | EN 14126:2003 (DSTU EN 14126:2008) |  |
| Period of validity, no less than | 4 years |  |
| PROTECTIVE GAS MASKS |
| Radiation protection | Protection against airborne radioactivity |  |
| Protection grade | 2-3 Class as per standard DSTU EN 136:2003 (ДСТУ EN 136:2003 in Cyrillic) |  |
| Material of gas mask | Halobutyl, bromobutyl rubber or EPDM (ethylene propylene diene monomer) |  |
| Filter | Threaded joint diameter of 40 mm. |  |
| Type of protection | Chemical, radioactive and biological agents |  |
| Components | Gas mask, Filtering and absorbing element. |  |
| Front | Protection against chemical, radioactive and biological agents for 24 hours or greater, Resistance to respiration of no more than 250 Pa, Legibility of speech of 95% or greater |  |
| Weight, no more than | 0.6 kg |  |
| Period of validity, no less than | 10 years |  |

| OTHER REQUIREMENTS |
| --- |
| Certified for use | Certification of equipment in Ukraine and/or in EU countries in accordance with the Ukrainian Legislation in force |  |

# SYSTEM FOR AUTOMATED RADATION MONITORING OF THE ENVIRONMENT (INCLUDING GROUND GAMMA SURVEY)

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 1 set |  |
| TECHNICAL SPECIFICATIONS |
| Purpose | Radiological monitoring on broad territories: (a) Automated radiation monitoring of the environment (gamma radiation background), (b) Automated radiation monitoring of the radioactive contamination of soil (ground gamma survey), (c) Marking the map with gamma dose rate measurements. |  |
| Type | Measuring system mounted in/on the vehicle |  |
| Functions | - Measurement of gamma radiation dose rate in two positions: (i) measurement of ambient dose equivalent rate H*(10) (alternatively, personal dose equivalent rate Hp(10)), (ii) measurement of the ground contamination (in collimator), - Marking the map with measurements; - Indication of measured values and radiological maps at the operator’s workplace (on one of the PCs); - Visual and audible alarm at the operator’s workplace when thresholds of radiation level are exceeded; - Recording the measurements in the application software; - Self-diagnosis of the detection units. |  |
| Operation mode | Up to 24 h of continuous operation |  |
| Setup time of operating mode, no more than | 15 min |  |
| Equipment | The supply shall comprise the following items: Gamma radiation detection units – 2 pcs, Collimator(s) for measurements of ground gamma survey – to be defined by the tenderer, Data processing unit – 1 pc, Visual and audible alarm unit – 1 pc, Fixing elements to attach the detection units and collimators to the vehicle, Complete set of cables (including data and power cables, sockets, connectors, interface/power adapters, etc.), Specialized software for setting up the equipment and data acquisition. |  |
| Power supply | From mains of AC current, 220 ±10% V, 50±5 Hz |  |
| Waterproof | IP 54 or better for detection unit |  |
| Shockproof | Yes |  |
| Mean time between failures, no less than | 4 000 hours |  |
| Service life, no less than | 4 years |  |
| DETECTION UNIT |
| Ionizing radiation | Gamma |  |
| Detector type | Energy compensated Geiger-Muller counter |  |
| Number of detection units | Two |  |
| Energy range of detected gamma radiation | From 60 keV to 3 MeV or better |  |
| Measurement range of ambient dose equivalent rate Hp(10) H*(10) (alternatively, personal dose equivalent rate Hp(10)) | From 0.01 μSv/h to 10 Sv/h or better |  |
| Uncertainty of measuring ambient dose equivalent rate Hp(10) when calibrated for 137Cs, no more than | ± 30% |  |
| ENVIRONMENTAL REQUIREMENTS |
| Ambient temperatures | From -20 to+45 0C |  |
| Relative humidity | Up to 95% at ambient temperature of 350C and lower, without condensation of moisture |  |
| OTHER REQUIREMENTS |
| Certified for use | Certification of equipment in Ukraine and/or in EU countries in accordance with the Ukrainian Legislation in force |  |
| Primary metrological verification | Yes, in Ukraine |  |

# WEATHER STATION

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 1 set |  |
| TECHNICAL SPECIFICATIONS |
| Type | Small wireless weather station installed on the roof of the Mobile radiological vehicle; equipped with the optical rain gauge. |  |
| Functions | Weather station shall display the following: ambient temperature, dew point temperature relative humidity, atmospheric pressure, speed and direction of wind, real time, speed and amount of precipitations, heat index, location according to GPS. |  |
| Access to data | From the automated operator work place (PC) and from mobile devices |  |
| ENVIRONMENTAL REQUIREMENTS |
| Ambient temperatures | From minus 30 to +50 0C |  |
| Relative humidity | From 0 to 100% relative humidity |  |
| Power supply | From 5 to 30V DC current |  |
| OTHER REQUIREMENTS |
| Certified for use | Certification of equipment in Ukraine and/or in EU countries in accordance with the Ukrainian Legislation in force |  |

# ISOTOPE IDENTFIFICATION DEVICE (with gamma and neutron channels)

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 1. Isotope identification device - 2 pcs. 2. Calibration/reference radiation source (in case if there is no built-in source/LED for spectrometric route stabilization) – 1 pc. |  |
| TECHNICAL SPECIFICATIONS |
| Type | Handheld gamma spectrometer (isotope identification device). Spectroscopic personal radiation detector (SPDR) is not acceptable |  |
| Ionizing radiation | Gamma and neutron |  |
| Detector type | Spectrometric scintillation detector |  |
| Dimensions of detector crystal, no less than | (30 x 30 mm |  |
| Energy range of detected gamma radiation | From 25 keV to 3 MeV or better |  |
| Relative energy resolution for line of 662 keV (137Cs) | 7 % or better |  |
| Automatic identification of radionuclides | - Nuclear materials: 233U, 235U, 239Pu; - Medical radionuclides: 18F, 67Ga, 99mTc, 111In, 123I, 125I, 131I, 133Xe, 192Ir, 201Tl; - Naturally occurring radioactive materials: 40K, 226Ra, 232Th, 238U; - Industrial radionuclides: 57Co, 60Co, 133Ba, 137Cs, 192Ir, 226Ra, 241Am. |  |
| Documented compliance to ANSI standards | Yes, ANSI N42.34 |  |
| Measurement of individual dose equivalent rate Hp(10) | From 0.01 μSv/h to 0.5 mSv/h or better |  |
| Channels of amplitude gamma spectrum, no less than | 1024 |  |
| Built-in LED and temperature system for automatic calibration | Yes. In case if there are no built-in source/LED for spectrometric route stabilization, a separate calibration/reference source(s) for automated calibration shall be supplied |  |
| Duration of routine energy calibration, no more than | 20 min |  |
| Recording gamma spectra in a non-volatile memory | Yes |  |
| Functions | - Automatic identification of gamma-emitting radionuclides and radionuclide analysis; - Measurement of individual dose equivalent rate; - Indication of measured values on the built-in display; - Storing spectra in a non-volatile memory and transferring to the PC; - Radiation safety alarm; - Automatic energy stabilization; - Assistance to energy calibration of the spectrometric detector. |  |
| Power supply | From integrated accumulators |  |
| Time of continuous operation at gamma background (not more than 0.5μSv/h) under non-alarm conditions, no less than | 8 hours |  |
| Total weight of all the components, no more than | 4 kg |  |
| Physical dimensions | Compact, portable, suitable for hand-held use |  |
| Waterproof, no less than | IP54 |  |
| Shockproof | Yes |  |
| Mean time between failures, no more than | 4000 hours |  |
| Service life, no less than | 6 years |  |
| ENVIRONMENTAL REQUIREMENTS |
| Ambient temperatures | From -20 to+45 0C |  |
| Relative humidity | Up to 95% at ambient temperature of 350C and lower, without condensation of moisture |  |
| OTHER REQUIREMENTS |
| Certified for use | Certification of equipment in Ukraine and/or in EU countries in accordance with the Ukrainian Legislation in force |  |
| Primary metrological verification | Yes, in Ukraine |  |

# MULTI-PURPOSE SURVEY METER (with alpha, beta, gamma and neutron channels)

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 1. Multi-purpose survey meter (with alpha, beta and gamma channels) – 2 sets, 2. Neutron radiation detection unit – 1 pc, 3. Marinelli vessel – 1 pc, 4. Shielding - 1 pc, 5. Reference/calibration radiation source, if necessary – 1 pc. |  |
| TECHNICAL SPECIFICATIONS |
| Type | Handheld device |  |
| Detection channels | Alpha, Beta, Gamma and Neutron independent channels |  |
| Gamma detector type | Combination of scintillation detector + Geiger-Muller counter is recommended |  |
| Detectors | Replaceable radiation detectors attached to the carrying handle/telescopic arm |  |
| Functions | Search, localization and evaluation of RN materials: - Measurement of ambient dose equivalent and ambient dose equivalent rate H*(10) by gamma and neutrons (alternatively, personal dose equivalent and personal dose equivalent rate Hp(10)); - Measurement of flux density of alpha and beta emitters; - Measurement of specific/volumetric activity of radionuclides in samples (for one of the Multi-purpose survey meter to be supplied); - Natural and intrinsic background suppression; - Indication of measured values on the LCD display; - Audible signalling of each registered particle in the search mode; - Audible and visual alarm when thresholds are exceeded; - Adjustable thresholds for alarm; - Recording the measurements in the non-volatile memory. |  |
| Equipment (to be supplied for both survey meters) | The Multi-purpose survey meter shall be equipped with the following: set of detectors of alpha, beta and gamma radiation, data display and processing unit, telescopic arm (of 1 m long or more), rechargeable battery, battery charger, specialized software for data treatment and analysis, carrying case. |  |
| Additional equipment (to be supplied for one of the survey meters) | The Multi-purpose survey meter shall be equipped with the following: neutron radiation detection unit, Marinelli vessel, Shielding, Reference/calibration radiation source, if necessary |  |
| ALPHA MEASUREMENTS |
| Energy range of detected alpha radiation | From 4 MeV to 7 MeV or better |  |
| Measurement range of alpha flux density | From 10 to 1(105 part./(cm2min) or better |  |
| Relative measurement error, no more than | ±20% |  |
| BETA MEASUREMENTS |
| Energy range of detected beta radiation | From 155 keV to 3 MeV or better |  |
| Measurement range of beta flux density | From 10 to 1(105 part./(cm2min) or better |  |
| Relative measurement error, no more than | ±20% |  |
| GAMMA MEASUREMENTS |
| Energy range of detected gamma radiation | From 60 keV to 3 MeV or better |  |
| Measurement range of ambient dose equivalent H*(10) (alternatively, personal dose equivalent Hp(10)) | From 0.1 μSv to 10 Sv or better |  |
| Measurement range of ambient dose equivalent rate H*(10) (alternatively, personal dose equivalent rate Hp(10)) | From 0.1 μSv/h to 100 mSv/h or better |  |
| Uncertainty of measuring ambient dose equivalent rate H*(10) when calibrated for 137Cs, no more than | ± 35% |  |
| Measurement range of specific (volumetric) activity for 137Cs in the geometry of Marinelli vessel | From 100 to 105 Bq/kg (Bq/l) or better |  |
| NEUTRON MEASUREMENTS |
| Energy range of detected neutron radiation | From 0.025 eV (thermal) to 14 MeV or better |  |
| Measurement range of ambient dose equivalent H*(10) (alternatively, personal dose equivalent Hp(10)) | From 0.1 μSv to 10 Sv or better |  |
| Measurement range of ambient dose equivalent rate H*(10) (alternatively, personal dose equivalent rate Hp(10)) | From 0.1 μSv/h to 100 mSv/h or better |  |
| PERFORMANCE |
| Alarm indication | Visual and audible alarms when thresholds are exceeded. Haptic alarm is an advantage. |  |
| Power supply | From integrated rechargeable battery |  |
| Time of continuous operation at gamma background (not more than 0.5μSv/h) under non-alarm conditions, no less than | 8 hours |  |
| Weight, no more than | 1.5 kg |  |
| Physical dimensions | Compact, portable, suitable for handheld use. Comfortable carrying handle and attachment to clothing is an advantage. |  |
| Shockproof | Yes |  |
| Waterproof | Protection grade of main components: IP 5X or better - for gamma radiation detectors, IP 2X or better - for alpha and beta radiation detectors. |  |
| Mean time between failures, no less than | 4000 hours |  |
| Service life, no less than | 6 years |  |
| ENVIRONMENTAL REQUIREMENTS |
| Ambient temperatures | From -20 to+45 0C |  |
| Relative humidity | Up to 95% at ambient temperature of 350C and lower, without condensation of moisture |  |
| OTHER REQUIREMENTS |
| Certified for use | Certification of equipment in Ukraine and/or in EU countries in accordance with the Ukrainian Legislation in force |  |
| Primary metrological verification | Yes, in Ukraine |  |

# SET FOR COLLECTING SAMPLES

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 2 pcs |  |
| TECHNICAL SPECIFICATIONS |
| Functions | Collection and transportation of samples for laboratory analysis: - Air samples, - Liquid samples, - Solid/bulk samples (e.g., soil) |  |
| Components | The set should include the following or equivalent items: - pipettes for suspicious liquids; - containers for suspicious liquids; - swab and container for solvents; - container for hazardous powder substances; - set of plastic containers; - PTFE containers for liquid and solid substances; - hermetically sealed container for chemical and biological hazardous substances; - plastic case. |  |

# BACKPACK FOR RADIONUCLIDE IDENTIFICATION AND SOURCE LOCATION

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 2 sets |  |
| TECHNICAL SPECIFICATIONS |
| Type | Portable monitor of RN materials in backpack |  |
| Detection channels | Gamma, neutron independent channels |  |
| Gamma channel | Spectrometric scintillation detector |  |
| Neutron channel | Alternatives to 3He counters are an advantage |  |
| Functions | Search, localization and evaluation of RN materials: - Automatic identification of gamma-emitting radionuclides and radionuclide analysis; - Measurement of individual dose equivalent rate; - Indication of measured values on the built-in display; - Indication of the direction to the found RN materials (radiation sources); - GPS positioning; - Internal mapping; - Assistance to localization of multiple RN materials (radiation sources) via user interface (incl. the calculation of estimated location for radiation source(s); - Radiation safety alarm; - Recording the measurements in conjunction to the time and place in the non-volatile memory; - Automatic energy stabilization. |  |
| Energy range of detected gamma radiation | From 30 keV to 3 MeV or better |  |
| Energy range of detected neutron radiation | From 0.025 eV to 14 MeV or better |  |
| Measurement range of individual dose equivalent rate Hp(10) | From 0.01 μSv/h to 100 μSv/h or better |  |
| Uncertainty of measuring individual dose equivalent rate Hp(10) when calibrated for 137Cs, no more than | ± 30% |  |
| Channels of amplitude gamma spectrum, no less than | 1024 |  |
| Relative energy resolution for line of 662 keV (137Cs) | 7.5 % or better |  |
| Alarm indication | Visual and audible alarms when thresholds are exceeded. Haptic alarm is an advantage. |  |
| Power supply | From integrated accumulators |  |
| Battery and battery life | Battery life should be greater than 6 hours under no alarm conditions for instruments |  |
| Weight, no more than | 10 kg |  |
| Physical dimensions | Compact, suitable for use as a backpack. |  |
| Shockproof | Yes |  |
| Waterproof, no less than | IP56 |  |
| Mean time between failures, no less than | 4000 hours |  |
| Service life, no less than | 6 years |  |
| ENVIRONMENTAL REQUIREMENTS |
| Ambient temperatures | From -20 to+45 0C |  |
| Relative humidity | Up to 95% at ambient temperature of 350C and lower, without condensation of moisture |  |
| OTHER REQUIREMENTS |
| Certified for use | Certification of equipment in Ukraine and/or in EU countries in accordance with the Ukrainian Legislation in force |  |
| Primary metrological verification | Yes, in Ukraine |  |

# PERSONAL ELECTRONIC DOSIMETERS

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Manufacturer | — |  |
| Model | — |  |
| Scope of supply | 1. Personal electronic dosimeters – 6 pcs., 2. Spare parts for 3 years of operation (accumulator batteries) – 1 set (to be defined by the tenderer) |  |
| TECHNICAL SPECIFICATIONS |
| Energy range of registered gamma radiation | From 60 keV to 3.0 MeV or better |  |
| Measurement range of individual dose equivalent rate Hp(10) | From 0.1 µSv/h to 1 Sv/h or better |  |
| Measurement range of individual dose equivalent Hp(10) | From 1 µSv to 9.9 Sv or better |  |
| Weight, no more than | 0.15 kg |  |
| Dimensions | Compact, suitable to be worn on a belt or to be carried in a pocket for hands free operations |  |
| Alarm indication | Visual and audible alarms when thresholds are exceeded. Haptic alarm is an advantage. |  |
| Functions | - Indication of measured values on the built-in display; - Adjustable thresholds for alarm; - Recording the measurements in the non-volatile memory. |  |
| Power supply | From accumulator battery |  |
| Battery and battery life | Accumulator batteries shall provide more than 2000 hours of continuous operation |  |
| Shockproof | Yes |  |
| Waterproof, no less than | IP54 |  |
| Mean time between failures, no less than | 4000 hours |  |
| Service life, no less than | 6 years |  |
| ENVIRONMENTAL REQUIREMENTS |
| Ambient temperatures | From -20 to+45 0C |  |
| Relative humidity | Up to 95% at ambient temperature of 350C and lower, without condensation of moisture |  |
| OTHER REQUIREMENTS |
| Certified for use | Certification of equipment in Ukraine and/or in EU countries in accordance with the Ukrainian Legislation in force |  |
| Primary metrological verification | Yes, in Ukraine |  |

# AUXILARY EQUIPMENT FOR POST-DETECTION ACTIVITIES

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Scope of supply | 1. Set of lights (searchlight and a tower for it, hand lamp, emergency light, head light, etc.) – 1 set (to be defined by the tenderer), 2. Warning signs and safety tapes to mark the boundaries (areas) of contamination with toxic and radioactive substances – 1 set, 3. Devices for automatic placement of safety tapes and warning signs – 2 sets (one stationary and one portable sets) 4. Photo/video cameras to ensure gathering of evidences and incident investigation – 2 pcs. |  |

# VISBILITY OF NORWEGIAN FINANCING

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| Visibility of Norwegian financing | Any equipment delivered under the contract should be clearly identified and should have metallic plates or indelible labels containing the flag of Norway and the phrase “Provided with support from the Government of Norway” in Ukrainian and in English. |  |

# DOCUMENTATION

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| DOCUMENTATION |
| Technical specifications/ conditions | In Ukrainian |  |
| Conceptual design of Mobile radiological complex | In Ukrainian |  |
| Programme and procedures of factory acceptance test | In Ukrainian and English |  |
| Vehicle documents | In Ukrainian and English |  |
| Passports/logbooks for Mobile radiological complex | In Ukrainian Passport/logbook shall contain the following data: - Completeness, - Placement of equipment, - Electric connection diagram, - Fuel and special liquids consumption, - Intended service life, etc. |  |
| Passports/logbooks for technical means | In Ukrainian |  |
| User’s manuals for technical means | In Ukrainian |  |
| Documents attesting certification of equipment in Ukraine and/or in EU countries | In Ukrainian and English |  |
| Declarations of compliance for technical means as per Technical Regulations on “Legally regulated measuring means” | In Ukrainian and English |  |
| Certificates of primary metrological verification for measuring instruments (issued by an authorized metrological centre in Ukraine) | In Ukrainian |  |
| Safety data sheet for decontamination solutions | In Ukrainian |  |
| Training documentation | In Ukrainian |  |
| Transportation documentation | In Ukrainian and English |  |

# DELIVERY TERMS AND CONDITIONS

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| DELIVERY TERMS AND CONDITIONS |
| Terms of Delivery | DDP, Incoterms 2010 |  |
| Place of Delivery | Pr. Peremohy 58, 93120, Lisichansk, Luhansk region, Ukraine |  |
| Delivery Time | ≤ 150 calendar days after the date of contract signature |  |
| PRESERVATION AND PACKAGING |
| Packaging | Protection for transportation, handling and reliable storage without re-preservation within 1 year upon delivery. |  |

# TRAINING COURSE: OPERATION, MAINTENANCE AND REPAIR

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| TRAINING |
| TRAINING COURSE |
| Place of training (training room to be provided by the Recipient) | Lisichansk unit of Luhansk border detachment of SBGS: Pr. Peremohy 58, 93120, Lisichansk, Luhansk region, Ukraine |  |
| Duration of training course | 3 days |  |
| Language of training course | Ukrainian |  |
| Number of trainees, up to | 30 persons |  |
| Themes to be covered | - Operating principles of all the units of Mobile radiological complex, - Detailed analysis of all the units of the Mobile radiological complex; - Use and operation of all the units of the Mobile radiological complex, - Measurement techniques, - Techniques of decontamination and applicable detergents, - Periodic metrological verification of measuring instruments, - Maintenance, - Minor repairs. |  |
| Theoretical part duration, no less than | 1,5 days |  |
| Practical part duration, no less than | 1,5 days: practical training at Recipient’s training room |  |
| Instructional video | Optional |  |
| Verification | Test, written form |  |
| TRAINING COURSE DOCUMENTATION |
| One certificate per trainee | In Ukrainian and English |  |
| List of materials to provide per trainee | - Set of training materials; - User’s documentation; - Training materials on one CD/DVD. |  |
| Results of test after training should be delivered to the Contracting Authority and Recipient | Yes |  |

# WARRANTY AND POST-WARRANTY SERVICES

|  | Contracting Authority’s Requirements | Tenderer’s Offer |
| --- | --- | --- |
| WARRANTY AND POST-WARRANTY SERVICES |
| Duration of warranty period | ( 24 months |  |
| Place of warranty repairs and maintenance | Ukraine |  |
| Presence of official representative of the manufacturer or authorised service centre in Ukraine | Yes, the official representative of the manufacturer or authorised service centre shall be available in Ukraine for: Warranty and post-warranty maintenance and repair of vehicle, Warranty and post-warranty maintenance and repair of radiation monitoring instruments. |  |
| Technical support during warranty and post-warranty period | - Technical support by email or telephone in Ukrainian/English to solve any technical problems (software failure, anomalous behavior, minor improvements concerning process, functional capabilities of data processing, etc.) and rectify any system-disabled state. - Consultancy support in relation to the periodic metrological verification by email or telephone in Ukrainian/English. |  |
| After-sales service | Compulsory after-sales service to be provided under a separate agreement with the Recipient shall include the following: - Maintenance and post warranty repair of the Equipment on the territory of Ukraine; - Rapid provision of spare parts and consumables. |  |

# TIME SCHEDULE

| No. | Action | Period of completion | Tenderer’s Offer |
| --- | --- | --- | --- |
| 1. | Submission of documents: Technical specifications/conditions; Conceptual design of Mobile radiological complex. | Within 45 calendar days after the signing of Contract |  |
| 2. | Submission of documents: Programme and procedures of factory acceptance tests, User’s manuals for technical means. | 2 weeks before the scheduled factory acceptance testing |  |
| 3. | Submission of documents: Vehicle documents, Passport/logbook for the Mobile radiological complex, Passports/logbooks for technical means, User’s manuals for technical means, Documents attesting certification of equipment in Ukraine and/or in EU countries, Declarations of compliance for technical means as per Technical Regulations on “Legally regulated measuring means”, Certificates of primary metrological verification for measuring instruments, Safety data sheets for decontamination solutions, Transportation documentation. | 2 weeks before the scheduled shipment and supplied with equipment |  |
| 4. | Submission of documents: Training programme, Training course documentation. | 2 weeks before the scheduled training |  |
| 5. | Equipment delivery | Within 150 calendar days after the signing of Contract |  |
| 6. | Training of personnel | Within 20* calendar days after the delivery of equipment |  |

Note: *) The personnel training must be carried out not later than within 20* calendar days after the delivery of Mobile radiological complex. The actual dates of training shall be confirmed by the Recipient and Contracting Authority no later than ten (10) days before the training course.

Authorized person on behalf of the Tenderer:

Name: ____________________________________________

Title: ______________________________________________

Signature: ________________________________________

Date: ______________________________________________